SUPPLEMENTARY MATERIAL

Table 1S. Measurements of sodium readings in ppm of the two standard solutions in a flame photometer. Source: own elaboration.

| Theoretical concentration (ppm) | Experimental concentration (ppm)* | The experimental concentration of Standard DM-513KC-64-4 (ppm)* |
| --- | --- | --- |
| 0.00 | 0.001±0.0010 | 0.001±0.0010 |
| 0.02 | 0.015 ± 0.0014 | 0.022±0.0025 |
| 0.05 | 0.034 ± 0.0085 | 0.053±0.0036 |
| 0.10 | 0.087 ± 0.0074 | 0.113±0.0144 |
| 0.20 | 0.185 ± 0.0202 | 0.212±0.0335 |
| 0.30 | 0.253 ± 0.0259 | 0.281±0.0915 |
| 0.50 | 0.438 ± 0.0427 | 0.493±0.0815 |
| 0.80 | 0.760 ± 0.0337 | 0.792±0.0695 |
| 1.00 | 0.925 ± 0.0316 | 0.981±0.0583 |

* Mean ± standard deviation, n = 3

Table 2S. Analysis of Variance for the linear regression curve. Source: own elaboration.

| Regression statistics |  |
| --- | --- |
| Multiple correlation coefficient | 0.99923096 |
| Coefficient of determination R2 | 0.99846251 |
| R2 adjusted | 0.99824287 |
| Standard error | 0.01424515 |
| Observations | 9 |

| Analysis of variance |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Freedom degrees | Sum of squares | Mean squares | F | The critical value of F |
| Regression | 1 | 0.92247 | 0.92247 | 4545.87459 | 4.15E-11 |
| Residues | 7 | 0.00142 | 0.00020 |  |  |
| Total | 8 | 0.92389 |  |  |  |

|  | Coeffi-cients | Standard error | Statistics t | Proba-bility | Lower 95 % | Superior 95 % | Lower 95.0 % | Superior 95.0 % |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Interception | -0.0094 | 0.0066 | -1.4262 | 0.1969 | -0.0250 | 0.006192163 | -0.025012751 | 0.006192163 |
| Theoretical concentration in (ppm) | 0.9361 | 0.0139 | 67.4231 | 0.0000 | 0.9032 | 0.968887023 | 0.903229238 | 0.968887023 |

Table 3S. Two-sample t-test assuming equal variances. Source: own elaboration.

|  | The experimental concentration of the standard solution | The experimental concentration of the saline solution |
| --- | --- | --- |
| Mean (ppm) | 0.330077778 | 0.299488889 |
| Variance (ppm) | 0.124216864 | 0.115486114 |
| Observations | 9 | 9 |
| Variance grouped | 0.119851489 |  |
| Hypothetical difference of means | 0 |  |
| Freedom degrees | 16 |  |
| Statistic t | 0.187433943 |  |
| P(T<=t) one-tailed | 0.426838348 |  |
| Critical value of t (one-tailed) | 1.745883676 |  |
| P(T<=t) two-tailed | 0.853676697 |  |
| Critical value of t (two-tailed) | 2.119905299 |  |